## Fill in
Medical Device Complaint and CAPA Record Template
qhsetech.com/templates/medical-device-complaint-capa-record-template ¦ General template, not legal or certification advice. Adapt it to your own procedures and to the requirements that apply to you. ¦ Free template by QHSEtech, CC BY 4.0: reuse and adapt it with credit (see the About sheet)
Field | Entry | Why it matters
Complaint number and date received |  | Section 820.35(a)(2) asks for the date the complaint was received. For a 30-day report a manufacturer is considered aware when any of its employees becomes aware of a reportable event (21 CFR 803.3(b)(2)), so the date that counts is the day the information first reached anyone in the organisation.
How received and received by |  | FDA's reporting page describes complaints as written, electronic or oral. Recording the channel shows which routes (service visits, distributors, returned devices) deliver information late and who needs to know how to pass it on.
Complainant name, address and phone |  | Section 820.35(a)(4) lists the complainant's name, address and phone number. They are needed for follow-up questions and for the reply that 820.35(a)(7) records.
Country and date of the event |  | The country shows which markets' rules section 6 has to check. The event date, which can be weeks before the complaint, is what the investigation and any report describe.
Device name, model and software version |  | Section 820.35(a)(1) asks for the name of the device. The model and the software or firmware version tie the complaint to a design and are what a search for similar complaints uses.
UDI or UPC, lot and serial number |  | Section 820.35(a)(3) asks for any UDI or UPC and other identification, and 820.35(c) for the UDI of each device or batch. The lot and serial number show which other devices may share the problem and where they went.
Device or sample returned and held |  | A returned device is often the best evidence of the cause. After a serious incident has been reported in the EU, Article 89(1) rules out investigations that alter the device or a batch sample in a way that may affect later evaluation of the cause before the competent authorities are informed.
Nature and details of the complaint |  | Section 820.35(a)(5) asks for the nature and details. Writing down what was reported, apart from what the investigation later finds, keeps the record clear about what was known on day 0.
Outcome for the patient or user |  | Death and serious injury drive the US questions in 803.50, and death, serious deterioration of health and serious public health threats drive the EU ones in Article 87. Unknown is an answer that calls for follow-up, not a no.
Is it a complaint? |  | FDA's reporting page describes a complaint as a communication alleging a deficiency in a device's identity, quality, durability, reliability, safety, effectiveness or performance after release for distribution. Under the QMSR the definitions in ISO 13485 apply (21 CFR 820.3), so the test is the standard's definition as your procedure sets it out. A No needs a reason and a note of where the matter went.
Possible failure to meet a specification |  | Section 820.35(a) asks for records of the review, evaluation and investigation of every complaint that involves a possible failure of the device, its labeling or its packaging to meet a specification.
Investigation decision and rationale |  | Under 820.35(a) a new investigation is not needed where a similar complaint has already been investigated, but the justification is recorded. The form asks for the earlier complaint or CAPA number and why that investigation covers this device, lot and failure.
Investigation and findings |  | The device evaluation, the batch record, the labeling and the similar complaints are the usual evidence, and the risk management file shows whether the hazard was foreseen. Section 803.50(b)(3) makes the manufacturer responsible for investigating each event and evaluating its cause.
Date of awareness |  | The US medical device reports and the EU serious incident reports run from awareness: 21 CFR 803.3(b) says when a manufacturer is considered to have become aware, and Article 87 counts from when the manufacturer becomes aware. A supplemental report runs from receipt of the new information, a US correction or removal from the day it starts, and an EU field safety corrective action is reported before it starts. Writing day 0 once, with the due dates worked out from it, keeps the deadlines visible.
US reportability (21 CFR 803) |  | Deaths, serious injuries and malfunctions that would be likely to cause or contribute to one if they recurred are reported within 30 calendar days, and a 5-day report is due where remedial action is needed to prevent an unreasonable risk of substantial harm or FDA has asked in writing. A decision not to report under 803.20(c)(2) keeps the information the qualified person used, and every decision goes in the MDR event file (803.18).
US correction or removal (21 CFR 806) |  | A correction or removal started to reduce a risk to health is reported to FDA within 10 working days of starting it (806.10(b)). One that is not reportable is still recorded, with a justification reviewed by a designated person (806.20).
EU reportability (MDR Articles 87 and 89) |  | A serious incident is reported immediately once a causal relationship is established or reasonably possible, and not later than 15 days after awareness; 10 days for a death or an unanticipated serious deterioration of health; 2 days for a serious public health threat. A manufacturer that is unsure still reports within the period (Article 87(7)), and a field safety corrective action is reported before it is undertaken, except in urgent cases (Article 87(8)).
CAPA and NCR link |  | The complaint record holds the complaint, the reportability decisions and the reply; the containment, disposition and why-chain sit in the nonconformance report and the CAPA, which ISO 13485 covers in clause 8.5.2 on corrective action. Linking them by number avoids two versions of the same analysis.
Root cause |  | One or two sentences that a reader of the complaint file alone can follow, with the reference to the full analysis. The cause category lets complaints be counted by design, manufacturing, supplier, labeling, packaging or use.
Corrections and corrective actions |  | Section 820.35(a)(6) asks for any correction or corrective action taken. The type column separates fixing this unit or lot from removing the cause, and a field action is checked against Part 806 and Article 87(1)(b).
Effectiveness check |  | The criterion is fixed before anyone looks at the data, so the check is able to fail. When it fails, the CAPA is reopened and the complaint stays open; the method box says which data (complaint trend, tests, audits) was used.
Reply to the complainant |  | Section 820.35(a)(7) asks for any reply to the complainant. The date, the channel and a summary show what the complainant was told and when; where there is no reply, the reason is recorded.
Closure and sign-off |  | Four roles sign: the complaint handler, the investigator, a regulatory reviewer for the reportability decisions and a quality approver. Section 803.18(e) allows the MDR event file to sit in the complaint file if it is prominently identified. The two retention periods printed are the US ones the texts state.

## Example (illustrative)
Illustrative example: invented, describes no real organisation
Cuff closure failure on an upper-arm blood pressure monitor (invented manufacturer, device, clinic, numbers and decisions)
Section | Entry
1 Intake | Complaint C-2026-0142, received by phone at customer service on 3 March 2026 and entered the same day. Complainant: a practice nurse at a general practice clinic in Ireland (name, address and phone recorded). Event first noticed on 27 February 2026.
2 Device | Upper-arm blood pressure monitor BPM-4 with reusable cuff C-22, firmware 2.1. UDI from the cuff label, lot L2511-118, cuff made in November 2025. Lot distributed in the US and the EU. One cuff, returned on 9 March and held in quarantine, untested, until evaluation.
3 What happened | The cuff's hook-and-loop closure opened during inflation twice in one week; the monitor stopped, showed error E3 and displayed no reading. The patient was not hurt and was measured with another cuff. Outcome: no injury. Immediate action: a replacement cuff from another lot sent on 4 March.
4 Is it a complaint? | Yes: it alleges a durability deficiency after release. Possible failure to meet the cuff's closure-strength specification: yes. Could it be reportable: unsure, so section 6 was started with 3 March as day 0.
5 Investigation | Investigated: no earlier investigation covered closure failures. Returned cuff: closure strength below the specification. Batch records: the fastener supplier changed its loop material in October 2025 under a change notice nobody assessed. Similar complaints: three more about closures, lots L2511-101 to L2512-130, in the complaint file from 1 January to 3 March. Risk file: closure failure listed, with the error message as its control.
6 Reportability | US: no death or serious injury. For the malfunction, the company's biomedical engineer concluded that a recurrence would not be likely to cause or contribute to a death or serious injury, because the monitor stops and shows an error instead of a wrong reading; not reported, with the information used filed in the MDR event file on 5 March. EU: an incident but, in the manufacturer's assessment, not a serious incident; not reported on its own and added to the trend data. No correction or removal in the field, so Part 806 and Article 87(1)(b) were not triggered; to be reassessed if a field action is proposed.
7 CAPA link | CAPA-2026-009 opened on 10 March and NCR-2026-031 raised for the affected fastener lots. Root cause: an unassessed supplier material change reduced closure strength. Cause category: supplier. Risk file updated with the new cause on 20 March.
8 Actions | 1 Replacement cuff to the clinic (correction; customer service; done 4 March). 2 Supplier returns to the earlier loop material (corrective action; supplier quality; due 15 April, done 10 April). 3 Closure-strength test on every incoming fastener lot (incoming inspection; done 20 March). 4 Supplier agreement: material changes need approval before shipment (purchasing; done 30 April).
9 Effectiveness | Criterion written on 20 March: no closure complaint on cuffs made from fastener lots received after 10 April in their first 90 days of distribution, and every incoming lot passes the test. Result on 31 July: no such complaint; six lots tested and passed. Effective: yes.
10 Reply | Phone call and email on 4 March saying a replacement was on its way; letter on 5 May with the findings and the change made.
11 Closure | Signed by the complaint handler, the investigator, the regulatory affairs reviewer and the quality manager; closed on 10 August 2026. The record and its MDR event file are in the electronic complaint file, with the event file marked as such.

## About
Template | Medical Device Complaint and CAPA Record Template
Printable form | /downloads/templates/medical-device-complaint-capa-record-template.pdf
Page | https://qhsetech.com/templates/medical-device-complaint-capa-record-template
Last updated | 2026-10-21
Note | General template, not legal or certification advice. Adapt it to your own procedures and to the requirements that apply to you.
What this workbook is | A spreadsheet for typing and for keeping a log. The PDF is the printable form. This file adds no formula and no rule.
Licence and source | Creative Commons Attribution 4.0 International (CC BY 4.0), https://creativecommons.org/licenses/by/4.0/. You may copy, adapt and share this file, also commercially, if you give credit, link to the licence and say if you changed it. Credit line: Medical Device Complaint and CAPA Record Template by QHSEtech (https://qhsetech.com/templates/medical-device-complaint-capa-record-template), licensed under CC BY 4.0 (https://creativecommons.org/licenses/by/4.0/). Template page: https://qhsetech.com/templates/medical-device-complaint-capa-record-template